ZXY600x Series DC-DC Power

User Manual

Rev1.1

2012.2.20

1. Introduction

ZXY600x series DC-DC Power Is our most recently developed with constant pressure and constant flow function of power supply products.Based on the structure of the BUCK switching power supply technology,Work frequency at 150KHz,With very wide input voltage range: DC15V to DC62V,Output voltage 0V to 60V Continuously adjusts(attention:BUCK can't Liters voltage,Output is always better than low input),Output current model: 3A,5A. Product Little calorific value, can be in calibration under current maximum continuous work under the full load.

The power output by microcomputer control,LCD can display:Output voltage,Output current,Output power,Ann hours,Output times and CV,CC status.And has kinds of intelligent shut off the function.Suitable for laboratory,The factory,Work shop ,Battery charging.

2. Each part function introduced

3. Features

3.1 Input voltage range: DC15~DC62V

3.2 Output voltage range:DC0~DC60V

3.3 Output current adjustment range:0A~3A(zxy6003),0A~5A(zxy6005)

3.4 Power effect ion:CV Less than 0.5%+10mV  CC Less than 1%+5mA

3.5 Load effect ion:CV Less than 0.5%+10mV  CC Less than 1%+5mA

3.6 Output ripple: <  50mVpp

3.7 100Hz wave transmission ratio：Less than 1/10000

3.8 Typical efficiency:Input voltage 32V,Output voltage 24V is 90%

3.9 Display precision:10mV,1mA

3.10 Voltage display error :±0.5%+2 bytes

3.11 Current display error：±1%+5 bytes

3.12 Protection types:

Over-voltage protection (OVP)：Set range 0.01V to 62.00V

Over-voltage protection (OCP)：Set range 0.001A to 5.100A

Over-power protection (OPP)：Set range 0.001W to 300.0W

Over-charging protection(OAH):Set range 1mAH to 900AH

3.13 Timeout protection(OFT)：Set range 1s to 90hours

4. Operation

Attention:4.1 to 4.5 are general operation,4.6 to 4.11 are special function set.

4.1 Correct wiring output,Plus-n-Minus can't wiring wrong,when open the power,the LCD default display preinstall voltage current,output is shut off,adjust voltage and current variable resistor may change voltage and current preinstall value,Figure1 is open power LCD display.

Figure1 Open power LCD display

Figure1 show voltage set 12.34V,current set1.235A，NOR equal to Normal,is

Normal, OFF show now power is shut off .

4.2 Now you can press <OUT> key,Power is work,when working LCD display real-time voltage,real-time current,real-time power,real-time charge capacity,

Figure2 Show power output voltage current state.

Voltage is 12.36V,this different to preinstall value with a small amount is normal,because no load,so current is 0A,and power is 0W，charge capacity

Is 0mAH,charge capacity is accumulated value,is have current,this value can increase.CV is eternal voltage model ,ON is power are output voltage,LCD right LED,CV and ON is light.

4.3 When have load,LCD display voltage and current value are on Figure3 show have load voltage and current value.

This figure show after have load,current value is 0.834A,Power output

is 10.30W,20mAH this value can all increasing,Now presskey，mAH can clear counter to accumulate.

4.4 Presskey,may change to time display,Figure4 show

Time format(Hour:Minute:Seconds),only when power is output voltage time can display,when output voltage shut off time stop, at the same time,at this state,presskey can clear time.

4.5 Presskey,return power model and mAH display model,and reduce the load,current can increase,when increase to 1.235A,output model is eternal current model,Figure5 show

This current value different to set value 1.235A have a little error,this is normal.state display CC is eternal current,at the same time LCD right "CC LED" and "ON LED "is light.

4.6 At the shut off state.pressandkey,can cycle output every function,press OK key in this function:

SET  U-CAL:calibration Voltage measure value

SET  I-CAL:calibration Current measure value

SET  OVP:Set Protection value

SET  OCP:Set the Over-Current protection value

SET  OPP:Set Power protection value

SET  OAH:Set Charging capacity protection value

SET  OFFTIME:Set timeout protection value

---SAVE  DATA！---：Set save parameters

Start up:"OFF"Set the boot is on,default is shut off

System Recover: System initialization,this function can initialization

Sound Enable:"ON"Set sound,default sound reminders is on.

SET "U-CAL" and SET "I-CAL"this function is adjust output voltage and current measure value,When you discover that measured value and the actual value have difference,you can use this two function to adjust,but this not suggest user operation,if you need,Please contact us.

4.7 OVP,OCP,OPP,OAH,OFFTIME this five is auto shut off,you can set voltage overrun,current overrun auto shut off,only by OCP for an example,the other all the same.

(a)At the tine output is shut off,pressandkey, LCD out Figure6:

(b)After,press <OK> key,set OCP protection, LCD out Figure7:

For this you can set OCP value,default is 0,means not use OCP function,only need set not is 0,can use OCP function,press <up> or <down> key may change this value,long press can quickly change, for example set 0.5A,this Figure8 show

(c)Press <OK> key confirm set,now it back to voltage current preinstall,

Press <OUT> key,and increase current,when it is 0.5A,auto shut off,and tip this shut off because over current OCP protection,for Figure9 show.

(d)So far, OCP function is working,but this set value is not save to FLASH,When power lost,this value can lost.

4.8 Save parameters function may save current set parameter save to FLASH,Next time automatic load at power on time.for Figure10 show.

Now press <OK> key may save parameters.

4.9 Default when power on output is shut off,for Figure11,

press <OK> key, set "ON", when power on output not shut off.

4.10 if need cancel sound or use sound,for figure12.Press <OK> to set.

4.11 If you want to recovery factory set for figure13 show.

Now you only need press <OK> key can recovery factory value.

5. List

5.1 ZXY6005 host machine						1PCS

5.2 User Manual 					       	1PCS