
SDS1000X-E Disassembly Instructions
Security Consideration 

Only qualified personnel should perform the disassembly procedures. Disconnect the power before removing DSO cover. Otherwise, personal injuries or damages to the components may occur.

Avoid Electric Shock Hazardous voltages exist on the LCD module and power supply module. To avoid electrical shock, disconnect the power cord from the instrument, and then wait at least three minutes for the capacitors in the instrument to discharge before beginning the disassembly.

Preventing ESD Almost all electrical components can be damaged by electrostatic discharge (ESD) during handling. Component damages can occur at electrostatic discharge voltages as low as 50 volts. The following guidelines will help preventing ESD damage when servicing the instrument or any electronic device.

· Disassemble instruments only in a static-free work area. 

· Use a conductive work area to reduce static charges. 

· Use a conductive wrist strap to reduce static charge accumulation. 

· Minimize handling. 

· Keep replacement parts in original static-free packaging. 

· Remove all plastic, foam, vinyl, paper and other static-generating materials from the immediate work area. 

· Use only anti-static solder suckers. 
Required Tools

· Multifunction screwdriver
· Antistatic gloves

· 16mm or 5/8” AF deep socket wrench For BNC nuts/fasteners
Disassembly Procedures

1. Remove the knobs. Remove each front-panel knob by firmly grasping the knob (with pliers protected by a soft cloth to prevent scratches, if necessary) and pulling it away from the front panel.
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2. Remove the 2 pcs PM3*10 screws between the front and rear shells and the 2 pcs KM3*6 screws below the handle, then remove the rear shell from the instrument.
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3. Remove the 4pcs KM3*6 screws located on the metal shell, and then remove the rear metal cover and the upper metal cover. 
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4. Remove the 4 pcs KM3*32 screws fixing the fan, and disconnect the cable from the fan to the main board, then remove the fan; Remove the 2 pcs M4*8 screws between the AC inlet and the metal frame; Remove the 4 pcs PWM3*8 screws between the AC-DC board and the metal frame; Disconnect the grounding plug from the AC-DC board to the mental frame, disconnect the cable from the AC-DC board to the main board, and then remove the AC-DC board; 
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5. Remove the 2 pcs PM3*6 screws fixing the insulation sheet, then remove the insulation sheet.
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6. Remove the 6 pcs PC3*8 screws between the metal frame and the front module, and dicconnect the front panel cable and the display cable from the main board, then the front module can be separated from the metal frame.
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7. Remove the 3 pcs nuts from the BNC connectors.
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8. Remove the 5 pcs PM3*6 screws on the main board, then the main board can be separated from the metal frame.
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